## Slide 1: A collaborative project between
Metrosol Limited & NPL
A Practical Johnson Noise
Thermometer

## Slide 2
 There is a demand for better thermometers
 Current thermometers drift, especially in harsh environments
Why and what’s the solution?
The Need for a Better Thermometer

## Slide 3: Current (secondary) Thermometers
Measure property affected by temperature
Calibration
Temperature scale
ITS-90
True thermodynamic temperature
Things other than
temperature can affect
this property

## Slide 4: Johnson Noise (primary) Thermometer
Measure Johnson Noise
Fundamental
Equation
True thermodynamic temperature

## Slide 5
Using Johnson Noise to Make a Thermometer
RMS

## Slide 6
An innovation project by ……
The Johnson-Nyquist Equation
= the RMS voltage
= Botzmann’s constant
= the thermodynamic temperature of the resistor
= the resistance of the resistor
= the measurement bandwidth
Problems:
extremely small (at limits of measurement)
hard to define/determine

## Slide 7
An innovation project by ……
The Johnson-Nyquist Equation
= the RMS voltage
= Botzmann’s constant
= the thermodynamic temperature of the resistor
= the resistance of the resistor
= the measurement bandwidth
Problems:
extremely small (at limits of measurement)
hard to define/determine

## Slide 8
An innovation project by ……
Amplifiers A & B amplify the same signal from the sense resistor but add their own noise
The correlator supresses uncorrelated noise and detects only the signal common to both channels i.e. the sensor’s Johnson noise
This allows us to measure the extremely small Johnson noise signal provided the correlation time is sufficiently long
Prior Art Johnson Noise Thermometer

## Slide 9
An innovation project by ……
The amplifiers switch between measuring the Johnson noise and a known calibration signal (a “white noise” source of known power density)
The ratio of the measurements allows the Johnson noise power spectral density to be determined provided is constant
Prior Art Johnson Noise Thermometer

## Slide 10
An innovation project by ……
The switching between the 2 measurements means that changes, mainly due to the capacitance of the cables
Solution has been to match the characteristics of the transmission lines carefully
Even with short cables, system is limited to around 100Ω & 100kHz
Prior Art Johnson Noise Thermometer

## Slide 11
An innovation project by ……
Calibration signal injected as current, which generates a calibration voltage in the sensor
The Johnson noise & calibration signal experience the same frequency response
The signals are subsequently separated completely using digital signal processing
Metrosol Johnson Noise Thermometer

## Slide 12
Pseudo Random Noise Calibration Signal

## Slide 13
Pseudo Random Noise Calibration Signal

## Slide 14
Pseudo Random Noise Calibration Signal

## Slide 15
Pseudo Random Noise Calibration Signal

## Slide 16
Results = Linearity & Offset

## Slide 17
Results: Accuracy
Chart: Time Series

## Slide 18
Prior art: high cost, room sized experiments operating in a screened room:
Prior Art JNT (at NIST)

## Slide 19
Proof of principal prototype, working in normal laboratory environment:
First “Practical” Johnson Noise Thermometer

## Slide 20
JNT1: 35kg, 70L & -20 to +120°C to…
JNT2: 1kg, 1L & -100 to 1,000°C:
Second Stage Prototype “JNT2”

## Slide 21
Progress So Far

## Slide 22
Create PRN data file
Quantise to 16-bits (1 in 65536 = 0.0015%)
Run data through FFT + FD correlator
Calculate:
  Ratio: “empty” FFT bins/tones bins
  Error in tone bins (peak & standard deviation)
Uncertainty Budget: waveform quantisation

## Slide 23
Correlator “eliminates” amplifier noise:
Initial Promising Results

## Slide 24
Correlator “eliminates” amplifier noise:
Initial Promising Results

## Slide 25
3 Year Project: Oct 2017 to Sep 2020
Technical progress on track
Starting to engage with market
Metrology partner:
Civil Nuclear partner:
Project Progress

## Slide 26: A collaborative project between
Metrosol Limited & NPL
For more information, visit….
www.johnson-noise-thermometer.com
www.facebook.com/johnsonnoisethermometer/