## Slide 1: 2nd EMPRESS Workshop 18-04-2018 | In situ temperature measurements by optical spectroscopy from lab to industrial scale
Senior Scientist Alexander Fateev (alfa@kt.dtu.dk)

## Slide 2
Background
Worse case scenario: large scale measurements (power plants and ship engines)
Best case scenario: lab scale (WP4 NPL standard burner)
Conclusions so far
2nd EMPRESS Workshop| Outline

## Slide 3: Background | Broad range of Applications…
Advanced, Fast, Precise measurements New insights on processes
Emission from surfaces Leak valve bottle Extractive probe boiler Exhaust air craft engine
Heat transfer,
energy balance
IR & video imaging
Gas composition,
temperature
Mixing, flows and properties
Bring university or research skills together
with practice

## Slide 4: Background | Temperature Laboratory
Temperature
Emissivity
External service
Customer problems
(plastic, light sources, sensors, heat transfer,…)
NMI, National Reference Laboratory for non-contact Temperature
State-of-art instruments
European cooperation, Courses for industry
High international standard of services
Range: -196C to 1600C accredited, largest Danish Temperature Lab.
ISO 17025

## Slide 5
Light source (a lamp, hot gas)  Medium ((hot) gas)  Detector
Spectral range: 180 nm – 400 cm-1
Detector: spectrometer with grating (UV/VIS/IR) or FTIR (IR)
Background | Optical Spectroscopy

## Slide 6
Emission spectroscopy:
Background | Optical Spectroscopy
Absorption spectroscopy:
ABS = ln[I0n/In] = s(l) x CONC x L
Fine structure of ABS gives Tgas
EMS = In= Lb x (1 – e[-s(l) x CONC x L])
If s(l) x CONC x L >> 1 then EMS  Lb and Tgas can be found from CO2 band radiance saturation at 2340 cm-1
where
Lb = Planck's function (black body curve)
s(l) = absorption cross section
CONC = concentration
L = pathlenght

## Slide 7: Complexity | Large Scale Measurements
Complexity:
get results first
trustful system
1500C is not uncommon
Expensive:
access possibilities
man power
time
Campaign at Block 7 Fynsværket (Denmark)

## Slide 8: Needs | Large Scale Measurements
IR image wood dust flame (video fuel mixing)
Boilers,
Flames (oil, gas, bio-masses),
VIS image grade flame (waste)
Engines (ships, jets),
Field campaigns (explosions)

## Slide 9: LSM | Temperature from IR measurements
Methods: Suction pyrometers (TC), FTIR (intensity), UV (shape features)
Burner region. Red: snap-shot,
Black: running average over 40 s

## Slide 10: LSM | Flame Gas Temperature

## Slide 11: LSM | Thermal Radiation Heat Transfer
Red: HITRAN 8.7% H2O at 1200C and over 15 m path Blue: Measured spectrum, Grey body 1186C and ε=0.562, Blackbody 1218C.
Spectral Intensity
Gas Temperature and major gas composition
Thermal radiation from gas
BB: 280 kW/m2
GB: 144 kW/m2
Thermal radiation from particles/walls
Total heat flux: Heat flux = εσT4

## Slide 12: LSM | Temperature from UV measurements
Spectrum of OH
UV head, boiler measurements
(Ø64 mm)
Miniature UV probe (Ø10 mm)

## Slide 13: LSM | Fast optical measurements
Imaging: Video and IR-camera (7 μs)
UV and IR Spectroscopy (1 ms)
Velocity, particle sizing (1 ns)
EU Hercules β: MAN Diesel 75% Load
Combine partner data with optical measurements with perfect synchronisation

## Slide 14: LSM | Fast gas concentration measurements (UV)
gas temperature
gas concentration
UV fiber optics for dynamic NO meas.

## Slide 15: LSM | Fast gas temperature measurements (IR)
Combination of
fast IR camera (640x512), InSb
grating IR spectrometer
3x optical IR fibers

3x simultaneous fast line-of-sight measurements (down to 7ms)

## Slide 16: LSM | Fast temperature measurements UV vs IR
Top
Middle
Bottom
very good agreement between IR(CO2 emission) and UV(NO absorption)
IR: less noisy data
wrong choice of TC (too slow)

## Slide 17
X
Y
Z
LSM | Best case scenario: lab scale in EMPRESS(WP4)
Optical set Up:
gas radiation optical measurements with a FTIR spectrometer;
Rx- single line-of-sight set Up purged by N2;
scanning in Rz and Ry-directions;
no CO2 self-absorption: full CO2(2340cm-1) band coverage (+ CO + H2O bands);
650-6000 cm-1 spectral range with linearized high-sensitivity MCT;
fast FTIR scanning, moderate spectral resolution (1cm-1 is enough, work with 4cm-1 and 8cm-1 as well);
spectra in W m-2 sr-1 (cm-1)-1 calibrated with portable BB at 796.1C (BB calibrated at DTU).
Modelling
use HITEMP2010 spectral database for CO2/H2O/CO modeling in 1500-2400 cm-1;
major isotopologues, Lorentz line shapes for single lines, cut off at ±8 cm-1 from line centers;
RHT emission spectra forward calculations with NPL’s Tgas profiles;
emission spectra calculations at selected Tgas and CO2 concentrations;
machine learning approach for Tgas profiles retrievals from single line-of-sight measurements tested.

## Slide 18
LSM | Emission spectra calculations
RHT forward calculations:
(MF Modest, Radiative Heat Transfer, 3rd , 2013)
CO2 example:
contribution from sub-columns in total emission

## Slide 19
Most difficult case (high CO concentration):
phi=1.4
no/with N2 co-flow
HAB=22mm and 32mm
experimental data (blue), modeling (olive)
Hyperspectral measurements underestimate H2O emission and show strong CO2 self-absorption from ambient air
no N2 co-flow, fit with NPL’s T-profile fit
N2 co-flow fit with NPL’s T-profile
N2 co-flow: UC3M data (red)
LSM | Some results in EMPRESS (WP4): IR

## Slide 20
HITEM2010 is very good for major part of CO2(-0.17%), H2O(0.63%) and CO bands (T2259K);
very good agreement with NPL’s T-profile;
no CO2 self-absorption;
experimental spectrum (blue) fitted with CO2(10.2%)+H2O(15.6%)+CO(1.17%) (red), phi=1, HAB=22mm;
HITEMP2010: quality of the database
Possibility to improve NPL’s temperature profiles:
original (blue) and corrected (olive) T-profiles at phi=1 HAB=20mm;
red: T-profile at HAB=10mm, phi=1.
LSM | Some results in EMPRESS (WP4): IR

## Slide 21
Optical spectroscopy is a powerful tool for in situ and on-line temperature measurements;
Broad application range: from lab to large scales;
No potentially upper limit for temperature measurements;
It can be used in various difficult environments;
2D and 3D measurement are possible;
Fast accurate measurements are possible at moderate equipment costs/size/weight;
In many cases gas composition can also be obtained;
Combination of modern spectral databases and powerful mathematical models (various temperature retrieval methods and machine learning approaches) give wide opportunities for new applications where temperature measurements are essential.
LSM&LSM | Conclusions so far

## Slide 22
Thank you for your attention