## Slide 1: WP4: Traceable combustion temperature measurement
Welcome to the National Physical Laboratory
EMPRESS 2nd Workshop 18th May 2018
NPL
Teddington
UK
Gavin Sutton NPL
Adam Greenen NPL
Juan Meléndez, Guillermo Andrés Guarnizo Herreño UC3M
Maria Jose Martin Hernandez, José Manuel Mantilla CEM
Alexander Fateev, Sønnik Clausen DTU
Paul Ewart, Anna-Lena Sahlberg UOX

## Slide 2
https://assets.publishing.service.gov.uk/government/uploads/system/uploads/attachment_data/file/633503/ECUK_2017.pdf
Energy Consumption in the UK July 2017
Why are combustion temperatures important?
WP4: Traceable Combustion Temperature Measurement
Combustion is the primary conversion process
By fuel type
By sector

## Slide 3
Why are combustion temperatures important?
WP4: Traceable Combustion Temperature Measurement
The temperature of any combustion
process critically affects:
  Chemical reaction rate
  Process efficiency
  Pollutant levels
  Product quality
  Rate of failure mechanisms
Achieving the correct temperature is critical
Qu: Why aren’t combustion temperatures routinely measured?
  Extremely harsh environment
  Probe survivability
  Flame evolution easily perturbed
  Access is challenging
  Model combustion process
  Validate model – test rig
    Scaled down
    Probe access
    Heavily instrumented
This is where we can help – good metrology: standards and instrumentation

## Slide 4
WP4: Traceable Combustion Temperature Measurement
Physical and optical probes
Example
Physical probes:
  Perturb the flame
  Give probe T not flame T
  Don’t last very long
  Access is challenging
  Slow response speed
Calibration source:
Region of hot gas of known T and composition traceable to ITS-90
Calibrate optical probe by measuring in this region
Optical probes:
  Emitted light can give T information
  Depends on ‘type’ of excitation
  Non-intrusive
  High response speed
  Access is still challenging
  Technically complex
  Needs calibration

## Slide 5
WP4: Traceable Combustion Temperature Measurement
WP4: Traceable Combustion Temperature Measurement
Objectives:
Develop portable STD flame with reproducible temperature (>1700 °C) and species concentrations
Using Laser Rayleigh scattering, determine temperature with < 1 % (x10 improvement)
Bring together three research groups working on different optical measurement techniques:
  Provide access to the STD flame
  Assess uncertainties
  Identifying sources of error
  Encourage collaboration and sharing of results to reduce uncertainties
At project end (Apr 18), STD flame facility available for calibration of third party optical techniques
Longer term - improve measurements in industry and academia
NPL STD flame under laser interrogation

## Slide 6
STD flame calibrated at NPL
Partner organisations – developing novel optical combustion thermometers
European project to enhance process efficiency through improved temperature measurement:
STD flame – circulated to partners
Comparison of techniques – publication
NPL facility available for others
WP4: Traceable Combustion Temperature Measurement
IR/UV spectroscopy
IR hyperspectral imaging
DFWM / LIGS
Portable STD Flame
Rayleigh scattering

## Slide 7
WP4: Traceable Combustion Temperature Measurement
The NPL standard flame:
  Hencken Burner – diffusion flamelets
  Propane / air flame –
  Low uncertainty flowmeters -
  Known post flame composition
  Traceable to ITS-90
  Portable
  Reproducible temperature -
Stoichiometric  fuel/air ratio for a balanced reaction (i.e. no excess oxygen)

## Slide 8
WP4: Traceable Combustion Temperature Measurement
The NPL standard flame - performance
Summary:
Temperatures lower than theoretical maximum due to heat loss to the burner surface
Reproducibility over 3 months  0.2 % of T
Suitable for use as a standard
a) Post-flame temperature: mean of 10 tests over 3 months
c) Post-flame temperature: Standard deviation of 10 tests over 3 months
b) Change in post-flame temperature relative to mean over 3 months

## Slide 9
WP4: Traceable Combustion Temperature Measurement
The NPL standard flame - performance
Example of temperature profile above the burner:  = 1.0
Green – HAB = 10 mm
Blue – HAB = 20 mm
Red – HAB = 30 mm
HAB = Height Above Burner
x
y
z
Laser
spot on burner showing
laser entry orientation

## Slide 10
WP4: Traceable Combustion Temperature Measurement
STD flame circulated to partners:
UC3M/CEM measurements completed May 17
UOX measurements completed Dec 17
DTU measurements: started Jan 18, ongoing (NPL will re-calibrate on return)
UC3M/CEM measurements – completed May 17
UOX measurements – completed Dec17
DTU measurements – started Jan 18

## Slide 11
WP4: Traceable Combustion Temperature Measurement
UC3M/CEM: Hyper-spectral imaging (1)
Measurement method applied to NPL standard flame presented at 18th International Metrology Congress (September 2017, Paris - France) - “Measurement of temperature and gas concentration in a flame using a FTIR imager”
CEM Blackbodies were used for radiance calibration of the instrument for flame temperatures up to 2400 ºC
Uncertainty in ε: 0.002 (k =2) → 0.12 ºC at 180 ºC and 0.26 ºC at 400 ºC (k = 2)
Final uncertainty of the black body temperature: 0.5 ºC
Ø 70 mm aperture blackbodies at 180 ºC and 400 ºC
Reference standard radiation thermometer
Hyperspectral Imager IR
Plane radiator

## Slide 12
WP4: Traceable Combustion Temperature Measurement
UC3M/CEM: Hyper-spectral imaging (2)
2 %
2 %

## Slide 13
WP4: Traceable Combustion Temperature Measurement
UC3M/CEM: Hyper-spectral imaging (3)
Following final measurements, system returned to NPL and accurate post-calibration verified by NPL
Example measurements of standard flame turned 90º (change of horizontal axis):

## Slide 14
WP4: Traceable Combustion Temperature Measurement
UC3M/CEM: Hyper-spectral imaging (4)
Completion of analysis of measurements and conclusions
Avoid flickering effect on standard flame with φ = 1,4 (rich configuration)
Avoid N2 co-flow – discrepancy between UC3M/NPL, source unknown
Measurement system and results will be presented in paper (lead by NPL)
Example data with flame flicker

## Slide 15
WP4: Traceable Combustion Temperature Measurement
UOX: Laser Induced Thermal Grating Scattering (LITGS) Thermometry in NO (1)
Preliminary measurements in methane/air slot burner: at l = 226 nm
Experimental arrangement for LITGS
Temperatures at 2 cm HAB:
Measurements consistent with adiabatic flame temperature
Burner

## Slide 16
WP4: Traceable Combustion Temperature Measurement
UOX: Laser Induced Thermal Grating Scattering (LITGS) Thermometry in NO (2)
LITGS thermometry in NO: NPL Standard Burner
Stoichiometric propane/air flame,
adiabatic flame temperature ~2280 K
Signal “distortion” by initial quenching process of LITGS at unknown rate.
Switch to Electrostrictive Grating
to avoid quenching effects.
Derived temperature exceeds
adiabatic flame temperature!

## Slide 17
WP4: Traceable Combustion Temperature Measurement
UOX: Laser Induced Thermal Grating Scattering (LITGS) Thermometry in NO (3)
Preliminary measurements in propane/air standard burner: at l = 1064 nm
Temperature vs position in flame
Weak, short-lived signals:
Large errors probably due to thermal grating contribution from H2O absorption?
High temperature, low pressure
environment is challenging for LIGS
τ=17.39 ns, T=2216 K
LIEGS signals (non-resonant)

## Slide 18
WP4: Traceable Combustion Temperature Measurement
DTU: Development of IR/UV spectroscopic flame thermometry systems (1)
IR measurements on NPL standard burner: set Up
Optical set Up:
Gas radiation optical measurements with a FTIR spectrometer
Rx- single line-of-sight set up purged by N2
Scanning in Rz and Ry directions
No CO2 self-absorption: full CO2(2340cm-1) band coverage (+ CO + H2O bands)
650-6000 cm-1 spectral range with linearized high-sensitivity MCT
Fast FTIR scanning, moderate spectral resolution (1cm-1 is enough, work with 4cm-1 and 8cm-1 as well)
Spectra calibrated with portable BB at 796.1 C (BB calibrated at DTU)
Modelling
HITEMP2010 spectral database for CO2/H2O/CO modeling in 1500-2400 cm-1
Major isotopologues, Lorentz line shapes for single lines, cut off at ±8 cm-1 from line centers
RHT emission spectra forward calculations with NPL’s Tgas profiles
Emission spectra calculations at selected Tgas and CO2 concentrations
Machine learning approach for Tgas profiles retrievals from single line-of-sight measurements tested
STD flame installed at DTU
Measuring emission spectra of STD flame

## Slide 19
WP4: Traceable Combustion Temperature Measurement
DTU: Development of IR/UV spectroscopic flame thermometry systems (2)
IR measurements on NPL standard burner: N2 co-flow issues
Most difficult case: phi=1.4 no/with N2 co-flow at HAB=22mm and 32mm. Current data in blue.
N2 co-flow: UC3M data (red) N2 co-flow: fit with NPL’s T-profile (olive) No co-flow, fit with NPL’s T-profile fit (olive)
Excessive N2 co-flow may cause cooling near edges of post-flame region
UC3M – N2 co-flow may cause issue outside flame region
DTU – N2 co-flow not an issue
Blue – measured
Green – model, based on NPL T profile
Blue – measured DTU
Red – measured UC3M
Blue – measured
Green – model, based on NPL T profile

## Slide 20
WP4: Traceable Combustion Temperature Measurement
DTU: Development of IR/UV spectroscopic flame thermometry systems (3)
HITEM2010 is very good for major part of CO2(-0.17%), H2O(0.63%) and CO bands (T2259K)
Very good agreement with NPL’s T-profile
No CO2 self-absorption
IR measurements on NPL standard burner: HITEMP2010 database quality
Experimental spectrum (blue) fitted with CO2(10.2%)+H2O(15.6%)+CO(1.17%) (red), phi=1, HAB=22mm
Original (blue) and corrected (olive) T-profiles at phi=1 HAB=20mm; red: T-profile at HAB=10mm, phi=1
DTU – improve NPL T profile by minimising differences in spectra

## Slide 21
WP4: Traceable Combustion Temperature Measurement
DTU: Development of IR/UV spectroscopic flame thermometry systems (4)
IR measurements on NPL standard burner
Key Results:
‘Clear’ flame emission spectra obtained at phi=0.8, 1 and 1.4 (no CO2 self absorption)
Excellent agreement between measured and calculated spectra with NPL’s T-profile data at phi=1 and 1.4 at HAB=10mm and 30mm, no N2 co-flow
NPL’s T-profile at HAB=20mm (phi=1, no N2 co-flow) has been corrected at far-flame boundary (Rx > 12mm) to fit better the measured spectra
Excellent agreement between measured and calculated spectra with NPL’s T-profile data at phi=1.4 and HAB=30mm with N2 co-flow
Excellent agreement between retrieved and calculated (GASEQ) CO2/H2O/CO concentrations at phi=1, HAB=20mm assuming 40 mm absorption pathlength (with/without N2 co-flow)
One spectrum fit [Tgas, CO2 concentration] is not enough to fit an experimentally measured spectrum in entire CO2 band - a full T-profile must be used
One fit [Tgas, CO2 concentration] gives only Tgas in the center
In “colder” post-flames (e.g. at higher HAB and/or phi=1.4) error in fit becomes larger if a single [Tgas, CO2 concentration] fit is used, i.e. emission from gas boundaries give more contribution to the total spectrum
Quality of the HITEMP2010 spectral database is very good for CO2/H2O/CO modeling in various spectral ranges
Machine learning approach for T-profile retrieval from single line-of-sight-measurements looks very promising, this will probably be used in coming EMPRESS2 project (collaboration with UC Merced, USA)

## Slide 22
WP4: Traceable Combustion Temperature Measurement
NPL: STD flame calibration following Trials at UC3M and UOX
No evidence of drift in the STD flame system
Final re-calibration required following return of system from DTU

## Slide 23
WP4: Traceable Combustion Temperature Measurement
Summary and outlook
  NPL:
  Portable STD flame developed
  Traceable Laser Rayleigh scattering thermometry system commissioned
  STD flame fully characterised – Ur (T)  0.5 %
  Peer review paper to be submitted: Evidence of 10 fold reduction in uncertainty
  UC3M:
  Demonstration of Hyperspectral imager in STD flame
  Good agreement with STD flame, low uncertainties achieved
  2D temperature and species profiles are possible
  Continue development in EMPRESS 2  industrial instrument
  DTU
  Measurement of IR emission spectra on STD flame
  Excellent agreement with NPL STD flame, low uncertainties achieved
  Validation of HITRAN 2010 database (CO2, H2O, CO)
  Improvement of STD flame T profile measurements
  Continue development in EMPRESS 2  industrial instrument
  UOX:
  LIGS demonstrated with flame generated NO (T  2300 K)
  Measurements on NPL STD flame were challenging
  Further development is needed for atmospheric pressure flames

## Slide 24
This project has received funding from the EMPIR programme co-financed by the Participating States and from the European Union’s Horizon 2020 research and innovation programme.
Thank you!